## Slide 1
Spectrophotometric Method Development and Validation for Application in Occupational Exposure Assessment of Trichloroethylene.
by
Lerato Monatisa
Intern Medical Scientist: MSIN 0006718
Analytical Services
NIOH Research day
30 November 2022
Healthy, Safe, Happy & Sustainable Workplaces
PROMOTING DECENT WORK THROUGH CUTTING EDGE RESEARCH, SPECIALISED SERVICES, INFORMATION, TEACHING AND TRAINING

## Slide 2
1
Selective method for TCE
2
Existing methods may be too contamination prone, too expensive, time consuming, or energy intensive
3
No method available
Need for the development of an alternative Method
INTRODUCTION

## Slide 3
INTRODUCTION
Trichloroethylene (TCE)
Grease from metal parts
Paint removers
Textile industries
Dry-cleaning facilities

## Slide 4: INTRODUCTION CONT’D
Exposure route:
Inhalation
Ingestion
Dermal Exposure
Biological Exposure Index (ACCIH,2019)
15 mg/L (114 mol/L)

## Slide 5
Acute
Chronic
LITERATURE REVIEW
Health effects of TCE
Fatigue 2,3,4
Headache2,3
Confusion 2,3,4
Blurred vision 2,3,4
Sleepiness 2,3,4
Central nervous system4
Reproductive system4
cardiovascular system
Pulmonary system4,5
Cancer4,5

## Slide 6
Analytical methods for the determination of TCE
LITERATURE REVIEW CONT’D
Author | Matrix | Sample preparation | Separation and Detection | LOD/LOQ (ng/ml)
Munch and Hautma., 1995 (EPA) | Water | LLE | GC-ECD | 1
Brown et al., 2003 | Tissues (livers, kidney, lungs) | LLE | GC-MS | 5
Liu et al.,2007 | blood, liver, lung, kidney and brain | HS-SPME | GC- MS | 0.25 and 0.75
Rastkari, N., Ahmadkhaniha, R., & Yunesian, M. (2012) | Urine | SPME | GC-MS | 0.005
Zhang, M., & Harrington, P. (2015) | Water | LLME–SPME | GC-MS | Not reported
G]]]g
GC: Gas Chromatography LLE: Liquid-Liquid Extraction
ECD: Electron Capture Detector HS-SPME: Head Space-Solid Phase Microextraction Phase
MS: Mass spectroscopy LLME: Liquid-Liquid Microextraction

## Slide 7
PROBLEM STATEMENT
The issues
TCE is a Toxic VOC: occupational Hazard
Current methods are expensive
The resolution
Employing spectrophotometric technique
An analytical method suitable for routine analysis
The outcome

## Slide 8: AIM AND OBJECTIVES
Aim
To develop and validate an analytical spectrophotometric method for application in occupational exposure assessment of trichloroethylene.
Objectives
To develop and validate a method for the routine analysis of TCE in human urine using UV-Vis spectrophotometry.
To increase the scope of tests offered in the Analytical Services department.

## Slide 9
Calibration standards (n=8)
50-700 µmol-L
Pyridine
5M NaOH
Ice cold water
Agilent Cary 60 UV-Vis spectrophotometer (λ =520nm)
Water bath (65°C)
Data analysis
Statistics: Excel, 2016
Glacial acetic acid
1% p-amino acetanilide
METHODOLOGY
Centrifuge (2500 RPM)

## Slide 10: RESULTS AND DISCUSSION
Figure 4: The effect of the temperature
Chart: Effect of temperature variations
| 22.07.2022 | 28.07.2022 | 02.08.2022 | 19.08.2022
QC1 | 86.3 | 96.9 | 97.3 | 87.9 | 98.5 | 100.5 | 86.7 | 99.2 | 98.2 | 101.7 | 96.3 | 106.8 | 100 | 94.5 | 108.5 | 99.5 | 95.1 | 108.9 | 98.6 | 93.9 | 105.4 | 94 | 93.7 | 80.5 | 93.9 | 94 | 93.7 | 80.5 | 93.9 | 95.9 | 90.9 | 81.5 | 86.1
QC2 | 325.3 | 305.2 | 296.6 | 318.2 | 312.9 | 328.8 | 321.1 | 301.1 | 307.9 | 263.1 | 303.4 | 287 | 261.7 | 250.5 | 300.1 | 261.9 | 251.2 | 299.9 | 250.5 | 300.4 | 300.9 | 252 | 304.2 | 263.2 | 311.7 | 252 | 304.2 | 263.2 | 312.1 | 320.2 | 246 | 291.2 | 257
Lab C30 Temperature ( ̊C) | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 21 | 23 | 23 | 23 | 23 | 23 | 23 | 23 | 23 | 22 | 22 | 22 | 22
Acceptable criteria:
QCs: (80-110%)
QC1 target: 100 µmol/L
QC2 target: 300 µmol/L
Temp: 15 – 25 °C

## Slide 11: RESULTS AND DISCUSSION CONT’D
Spectral characteristics and linearity
Figure 1: Absorption spectrum of TCE (50-700 µmol/L) at 520nm
Figure 2: TCE calibration curve
(λmax =520nm)
)

## Slide 12: RESULTS AND DISCUSSION CONT’D
Method Validation
Table 1: Analytical Figures of merit
Parameters | QC1 | QC2 | Acceptable Criteria
Recovery (%) | 95.62 | 96.60 | 80-110
Precision (% RSD) | 8.06 | 8.90 | 15 (max)
Analytical Figures of merit were calculated according to SOPs NIOH0391 and NIOH0371 as guided by TR 26 of SANAS
LOD: 0.16 µmol/L
LOQ: 0.54 µmol/L
UoM: 16.98%
LOD: Limit of detection
LOQ: Limit of quantification
UoM: Uncertainty of measurement

## Slide 13: RESULTS AND DISCUSSION CONT’D
Table 2: Comparison of TCE data from the present study with data from the other study
Author | Sample preparation | Separation and Detection | LOD (µmol/L)
Ye et al., 2017 | LLE | GC-ECD | 0.02
Monatisa, L. (2022) | LLE | Uv-VIS Spectrophometrer | 0.16
GC: Gas Chromatography
ECD: Electron Capture Detector
LLE:Liquid-Liquid Extraction
LOD: Limit of detection
Skilled personnel to operate the instrument and process the data.
Capital intensive
Not suitable for routine analysis

## Slide 14
1
A simple, sensitive and rapid method was developed and validated.
2
The analytical figures of merit were all within the acceptable ranges fulfilling SOPs NIOH0391 and NIOH0371.
3
Suitable for the routine analysis for biological exposure to TCE.
4
The method can also be used for research purposes.
CONCLUSION

## Slide 15: ACKNOWLEDGEMENT
Supervisors
Co-supervisors
Analytical Services

## Slide 16
WHO (2005) “Trichloroethylene in drinking-water,” World Health Organization .
Hansen, J. et al. (2013) “Risk of cancer among workers exposed to trichloroethylene: analysis of three Nordic cohort studies.,” Journal of the National Cancer Institute, 105(12). doi:10.1093/jnci/djt107.
Kumar, M. et al. (2009) “Study of genetic polymorphism in solvent exposed population and its correlation to in vitro effect of trichloroethylene on lymphocytes,” Journal of Environmental Biology, 30(5).
Singthong, S. et al. (2015) “Occupational health risks among trichloroethylene-exposed workers in a clock manufacturing factory,” Global journal of health science, 7(1). doi:10.5539/gjhs.v7n1p161.
REFERENCES

## Slide 17
Douglas, G.A., Ikeda, M. and Nogueira, D.P. (1981) “Recommended health-based limits in occupational exposure to selected organic solvents,” World Health Organization - Technical Report Series, No.664.
Brown, S. D., Dixon, A. M., Bruckner, J. v., & Bartlett, M. G. (2003). A validated GC-MS assay for the quantitation of trichloroethylene (TCE) from drinking water. International Journal of Environmental Analytical Chemistry, 83(5). https://doi.org/10.1080/0306731031000104713
Liu, Y., Xin, Q., Liu, D., Cai, Z., Zhang, Z., Lan, J., Wei, R., Li, D., & Hou, R. (2012). The application of HS-SPME-GC/MS in detecting chemical components in giant panda (Ailuropoda melanoleuca) urine. Acta Theriologica Sinica, 32(1).
Munch, D., & Hautman, D. (1995). Method 551.1: Determination of chlorination disinfection byproducts, chlorinated solvents , and halogenated pesticides/ herbicides in drinking water by liquid-liquid extraction and gas chromatography with electron-capture detection. Environmental Protection.
REFERENCES

## Slide 18
Rastkari, N., Ahmadkhaniha, R., & Yunesian, M. (2012). Simultaneous determination of trichloroethylene, perchloroethylene and trichloroacetic acid in human urine using solid-phase microextraction fibre coated with single-walled carbon nanotubes. International Journal of Environmental Analytical Chemistry, 92(14). https://doi.org/10.1080/03067319.2011.581367
Zhang, M., & Harrington, P. (2015). Determination of Trichloroethylene in Water by Liquid–Liquid Microextraction Assisted Solid Phase Microextraction. Chromatography, 2(1). https://doi.org/10.3390/chromatography2010066
REFERENCES

## Slide 19
“Your product is as good as the methods”
Ke ya Leboha